## Slide 1: Lecture 15: Nuclear Imaging: Introduction & Principle
Dr. Guohua Cao （曹国华） caogh@shanghaitech.edu.cn
School of Biomedical Engineering
ShanghaiTech University
2/16/2023

## Slide 2: Nuclear Imaging Overview
PET
SPECT

## Slide 3: Common Principle of Nuclear Imaging
Radioactive tracer is injected into subject
Uptake and distribution of tracer is influenced by subject physiology
Tracer undergoes radioactive decay and emits gamma rays
Gamma rays are detected by nuclear imaging camera and events stored
Data is reconstructed into image estimate of tracer distribution and uptake

## Slide 4: Example PET Image

## Slide 5: Electromagnetic Spectrum
Buildings
Humans
Butterflies
Needle Point Protozoans
Molecules
Atoms
Atomic Nuclei
10 4
10 8
1012
1015
1016
1018
1020
Radio
103
Microwave
10−2
Infrared
10−5
Visible
0.5×10 −6
Ultraviolet
10−8
X-ray
10−10
Gamma ray
10−12
Radiation Type Wavelength (m)
Approximate Scale of Wavelength
Frequency (Hz)
MRI
X-ray Nuclear Imaging Imaging
Optical Imaging

## Slide 6: Signal Sensitivity Comparison
Imaging Method | Sensitivity | 
Ultrasound | 10-3 Mol | 
CT | 10-3 Mol | 
MRI | 10-5 Mol | 
Bioluminescent | 10-8 Mol | 
Nuclear* | 10-9-10-12 | Mol

## Slide 7: Nuclear Imaging Source
radionuclide = atom with a specific number of neutrons and protons that undergoes radioactive decay
A X
Z N
A= mass number (#protons + neutrons=nucleons)
Z= atomic number (#protons)
N= neutron number (#neutrons)

## Slide 8: N N0  exp(t)
Radioactive Decay
Radioactive decay= mechanism by which an unstable nucleus is converted into a more stable one by emission of particles or photons
Half-life (τ1/2)= amount of time it takes a sample of radioactive atoms to halve
 = decay constant in inverse time

## Slide 9: Units of Activity
Bequerel (Bq)= SI unit defined in decays a second (dps)
Curie (Ci) = more commonly used for medical imaging
1 Ci = 37e10 Bq 1 mCi = 37 MBq
For nuclear imaging amounts of activity used are in the tens of µCi to tens of mCi
Particular activity amount used is determined by scanner capabilities, subject habitus, and tracer half-life

## Slide 10: n  p  e    energy
β- and (β-,γ) Decay
 
A A
Z X Z+1
A
 
A *  A
Z X  Z+1Y  Z+1Y
Examples
131I, 201Tl
Y

## Slide 11: p  n  e    energy
β+ and (β+,γ) Decay
 
A A
A
 
A *  A
Z X  Z-1Y
Z X  Z-1Y  Z-1Y
Annihilation Reaction
Positron and stationary electron produce two back-to-back 511 keV gamma rays (18F, 11C, 82Rb)

## Slide 12: p  e  n    energy
Electron Capture (EC) and (EC, γ) Decay
A
EC
X 
Z Z-1
A Y
Electron from closest orbits are usually the electrons captured (i.e. K or L shells)
Electron vacancy produces characteristic x-rays which can also be detected if of significant energy
Examples
67Ga

## Slide 13: Radionuclide Production
Generators
Cyclotron
Generator = Device that houses parent daughter radionuclide pair and allows extraction of the daughter
Cyclotron = Uses high power magnets to accelerate and collide very high energy charged particles into specialized targets

## Slide 14: Radiopharmaceutical Production
18F-Fluordeoxyglucose (FDG) synthesis
Two general labeling schemes
Direct substitution= Exchange of the stable atom with a radionuclide of the same chemical element
Analogs = Exchange of stable atom with a radionuclide of a
different chemical element (e.g. 18FDG)

## Slide 15
Radiopharmaceuticals for Medical Imaging
SPECT
PET

## Slide 16: Effective Half-life
Biological half-life (τbio)= rate that radiopharmaceutical is removed from the body
1  eff  1 1/2 1  bio
18F 1.9 h
11C 20 m
18FDG (metabolism)
11C-Choline
(receptor binding)
15O-H2O
(blood flow)
seconds
15O 2 m
Cell tracking
64Cu 12.7 h

## Slide 17: Radiopharmaceutical Preferred Qualities
Gamma ray emission in acceptable energy range
  50-600 keV for min subject and max detector attenuation
  Particulate and low energy gammas only add dose
Radionuclide and biological half-lives are similar
  Radionuclide half-life must be >= to biological process, but if too long will add dose to subject and minimize scan SNR
Total mass introduced to subject does not alter physiology
  Typically ng levels

## Slide 18: Interaction of Gamma Rays With Matter
Gamma rays interact with the subject and scanner mainly through the
photoelectric effect and Compton scattering
Compton scattering dominates µ at nuclear imaging energies!
µ for biological tissues is typically significantly lower at nuclear imaging energies c/w x-rays
Consequence
More material or material with higher µ/ρ is required in nuclear imaging to attenuate photons than x-ray imaging

## Slide 19: Attenuation Properties of Interest
80 keV
Ratio of totals
Mass Attenuation Coefficients
CT
500 keV ~PET
Material | Photoelec. | Compton | Total |  | Photoelec. | Compton | Total |  | 80 keV:500 keV
Air | 0.006 | 0.161 | 0.167 |  | ,0.001 | 0.087 | 0.087 |  | 1.92
Water | 0.006 | 0.178 | 0.184 |  | ,0.001 | 0.097 | 0.097 |  | 1.90
Muscle | 0.006 | 0.176 | 0.182 |  | ,0.001 | 0.096 | 0.096 |  | 1.90
Bone | 0.034 | 0.175 | 0.209 |  | ,0.001 | 0.093 | 0.093 |  | 2.26
Teflon | --- | --- | 0.168 |  | --- | --- | 0.087 |  | 1.93
Lead | - | - | 2.419 |  | - | - | 0.164 |  | 14.75
Tungsten | - | - | 7.810 |  | - | - | 0.138 |  | 56.52
Material |  | 80 keV |  | 500 keV | Ratio (500 keV:80 keV)
Water | 41.51 |  | 71.46 |  | 1.72
Bone | 17.27 |  | 38.82 |  | 2.25
Lead | 0.25 |  | 3.72 |  | 14.88
Half Value Layers (mm)

## Slide 20: Nuclear Counting Systems
What are the basic requirements necessary to efficiently count gamma rays?
Detect gamma rays
Distinguish energy of photons
Separate detection of photons in time
Signal Conditioning + Processing
ADC
Gamma Ray Detector
Computer
140 keV
320 keV

## Slide 21: Scintillation Detectors
Scintillator
Photodetector
Ionizing Radiation
Light
Scintillator= attenuates and converts incident radiation into light near the visible spectrum (both visible and UV)
Photodetector = Converts photons from scintillation light to an amplified electrical signal

## Slide 22: Scintillators
Organic: Scintillation is inherent property of the molecule (not used in nuclear imaging)
Inorganic: Ionic crystalline solids that convert energy into light by
nature of structure
Inorganic mechanism: Photoelectric and Compton interactions excite electrons which produce light upon return to ground state
Number of scintillation photons produced is proportional to absorbed gamma ray energy!

## Slide 23: Important Scintillator Properties
Attenuation Coefficient: Determines sensitivity of scintillator for absorbing high energy photons. Dependent on Z and density.
Light Yield: Efficiency of scintillator to convert absorbed γ energy to scintillation light. Influences energy separation of detector.
Decay Time: Speed of the scintillation pulse. Time for amplitude to return to 1/e
Property |  | NaI(Tl) |  | CsI(Tl) | LSO
Density (g/cm3) | 3.67 |  | 4.51 |  | 7.40
Effective Z | 50 |  | 54 |  | 66
Decay time (nsec) | 230 |  | 1000 |  | 40
Photon Yield (per keV) | 38 |  | 52 |  | 20-30

## Slide 24: Photodetectors
Purpose: Convert photons from scintillation light to an electrical signal
Photomultiplier Tubes (PMTs)

## Slide 25: Readout Electronics
Amplifier
ADC
PMT
High Voltage Power Source
Computer |  | Pulse-Height Analyzer
High Voltage Source: Relatively high (~1000 V) and stable (<1% variation) for PMTs
Amplifier: Amplifies low voltage signal from PMT and shapes pulses
Pulse-Height Analyzer: Histograms amplitude (energy) of pulses to produce energy spectrum

## Slide 26: Energy Spectra
Energy Resolution = (FWHM of photopeak)/(true energy)

## Slide 27: Gamma Camera Overview

## Slide 28: Collimator Function
Purpose: Projects image of tracer distribution onto scintillator by restricting the gamma ray incident angle range
Septa: Lead (high Z materials) walls surrounding holes that attenuate incident gamma rays
Projection on scintillator
True

## Slide 29: Collimator Function
Purpose: Projects image of tracer distribution onto scintillator by restricting the gamma ray incident angle range
Septa: Lead walls surrounding holes that attenuate incident gamma rays
Projection on scintillator
True
Measured

## Slide 30: Collimator Function and Photon Binning
2D Image Formed by Scintillator
Photons are binned in space based on position along the sctintillator face
Image is formed by the different counts in each histogram bin

## Slide 31: Collimator Types
Parallel Hole
Diverging
Converging
Pinhole
Different designs control magnification / minification, of object and FOV
Parallel hole collimator is most frequently used
Magnified Collimator View

## Slide 32
Parallel Hole Collimator Performance
k=constant related to collimator hole shape (5.8e-2 to 7.8e-2)
Geometric efficiency= fraction of γ-rays passing through collimator per a source γ-ray emission
Geometric efficiency for parallel hole collimator is independent of distance to subject (z)

## Slide 33: Parallel Hole Collimator Performance Tradeoffs
1. Collimator geometric efficiency and resolution are strongly linked
Improving Resolution
Reduction in Geometric Efficiency
Reduction in Image SNR
2. Geometry needs to be optimized for the energy of the γ-rays to be imaged
Increasing Energy
Reduction in Septal Material Attenuation
Requires thicker septa to attenuate γ-rays

## Slide 34
Parallel Hole Collimator Performance Examples
Different parallel hole collimator geometries are needed to image different combinations of radionuclides and clinical applications
Geometric

## Slide 35: Gamma Camera Scintillation Detector
Light Guide
PMTs
NaI(Tl) Crystal
NaI(Tl) Crystal = single continuous large area scintillator (25-50 cm diameter) and ~1 cm thick
Light Guide = spreads out and guides light to improve event positioning and avoid dead areas
PMTs = hexagonal array allows for positioning of γ-ray interactions in the NaI(Tl) crystal using Anger logic.
Side View Back View

## Slide 36
γ-ray
Anger Logic
Purpose: Calculate 2D γ-ray interaction point along face of crystal using signal amplitude from PMT array
Anger Logic Circuit
signal
X 
X  X  
X  X  
X Position for Logic Circuit

## Slide 37: Intrinsic Spatial Resolution
Accounts for effects of scintillator and electronics alone
Depends on two main factors:
Degree of multiple Compton scattering in scintillator
Statistical fluctuations of scintillator photons
γ-ray
T
Increasing γ-ray energy
Increase in scintillation photons
Improvement in intrinsic resolution
Increasing crystal T
Increase in multiple scatter events
Reduces intrinsic resolution

## Slide 38: Data Acquisition Electronics

## Slide 39: Energy Windowing
Energy Window
Photopeak
Compton Scatter
Scattered events carry little information about the tracer distribution (reduce image spatial resolution)
Selecting only events in a certain energy range (energy window) greatly reduces influence of scatter

## Slide 40: System Spatial Resolution
Influenced by collimator, intrinsic spatial resolutions, and Compton scattering in the subject
Rsystem 
gamma
R2
R
coll
2
R
2
Compton
Intrinsic
Collimator Compton
Collimator resolution has the largest impact on system resolution
Compton scatter mainly influences the tails of the PSF
Formula is only valid for a given distance from the collimator
Typically 8-14 mm for 99mTc